## Slide 1
Creativity & Confidence
www.atikorea.com
Lux
  Optical Chemical Concentration Analyzer

## Slide 2
Introduction : Lux
01
Application : SC-1
03
Application : DSP
02
Contents
Quality & Reliability
04
New functions
05

## Slide 3
Size(mm) : 369 x 309 x 540 (WxDxH)
Sensor module (UV + IR)
Principle
OAS : Optical Absorption Spectroscopy
Light Source : UV, IR
Features
In-line type / real-time monitoring
Auto validation / Auto calibration
Parts life-time monitoring
Repeatability : ±2% of display
Constant Temp. & Flow rate
VOC & Service quick response
Application
DSP chemical – Hydrogen Peroxide(H2O2), Sulfuric Acid(H2SO4)
SC-1 chemical - Hydrogen Peroxide(H2O2), Ammonia(NH3)
Etc. – Nitric Acid(HNO3)
Introduction: Lux

## Slide 4
 | Application | 
Target chemical | SC-1 (H2O2 + NH4OH) | DSP (H2O2 + H2SO4)
Target concentration | H2O2 : 5.43wt% NH4OH : 1.03wt% | H2O2 : 3.7wt% H2SO4 : 8.5wt%
Absorption wavelength | H2O2 : 280nm (UV) NH4OH : 2210nm (IR) | H2O2 : 280nm (UV) H2SO4 : 2200nm (IR)
H2O2 (in DSP+/SC-1) : UV - 280nm
H2SO4 (in DSP+) : IR - 2200nm
2200 nm
NH4OH (in SC-1) : IR - 2210nm
2210 nm
Chemical 고유 흡광 파장 광원 사용
HNO3 : UV - 280nm
280 nm
Introduction: Lux

## Slide 5
1. Calibration curve fitting
1.1 Test information
Chemical concentration range
  - H2O2 : 1/2/3/4/5 (wt%)
  - H2SO4 : 2/4/6/8/10 (wt%)
측정 시간 – 시료마다 5분
기준 spec. R2≥0.9985
1.2 평가결과
 | UV (ratio) | IR (ratio)
R2 | 0.99994 | 0.99999
Application: DSP
R²=0.99994
R²=0.99999

## Slide 6
 | Target conc. |  | Calculation conc. | 
 | H2SO4 | H2O2 | H2SO4 | H2O2
1-1 | 1.000 | 2.195 | 0.997 | 2.196
2-2 | 2.056 | 3.967 | 2.063 | 3.954
3-3 | 3.010 | 6.000 | 2.995 | 5.999
4-4 | 3.991 | 8.058 | 3.984 | 8.055
5-5 | 5.008 | 10.004 | 5.012 | 9.988
2. DSP 농도 변별력 test
2.1 Test information
2.2 평가결과
Target chemical
  - H2O2 : 1/2/3/4/5 (wt%)
  - H2SO4 : 2/4/6/8/10 (wt%)
측정 시간 – 시료마다 5분
목표 spec.
  - H2O2 (wt%) ±0.3%
  - H2SO4 (wt%) ±0.2%
- 시료 변경구간
Application: DSP

## Slide 7
1. Calibration curve fitting
1.1 Test information
Chemical concentration range
  - H2O2 : 0/2/4/6/8/10 (wt%)
  - NH4OH : 0/1/2/3/4/5 (wt%)
측정 시간 – 시료마다 5분
기준 spec. R2≥0.98
1.3 Test information
SC-1 농도 25종 calibration curve 작성
SC-1 chemical reaction : UV/IR signal interference
1.4 평가결과
1.2 평가결과
 | UV (ratio) | IR (ratio)
R2 | 0.99993 | 0.99996
Application: SC-1

## Slide 8
 | Target conc. |  | Calculation conc. | 
 | NH4OH | H2O2 | NH4OH | H2O2
1-1 | 0.7008 | 5.0004 | 0.7555 | 4.7479
3-3 | 0.9996 | 5.5008 | 1.0725 | 5.4164
5-5 | 1.3004 | 6.0007 | 1.3924 | 5.9415
1-1
3-3
5-5
1-1
3-3
5-5
2. SC-1 농도 변별력 test
2.1 Test information
2.2 평가결과
Target chemical
  - H2O2 : 5.0/5.5/6.0 (wt%)
  - NH4OH : 0.7/1.0/1.3 (wt%)
측정 시간 – 시료마다 5분
목표 spec.
  - H2O2 (wt%) ±0.3%
  - NH4OH (wt%) ±0.2%
Spec. in
  - H2O2 conc. error ≤0.3%
  - NH4OH conc. error ≤0.2%
- 시료 변경구간
Application: SC-1

## Slide 9
 | UV (ratio) | IR (ratio)
Average | 0.05074 | 2.14871
STDEV | 0.00004 | 0.00122
RSD(%) | 0.09 | 0.06
1차
2차
3차
1차
2차
3차
1. 재현성 test [SC-1]
1.1 Test information
1.2 평가결과
Target chemical: H2O2 (5.5 wt%), NH4OH (1.0 wt%)
측정 시간 – 시료마다 5분
측정 횟수 – 3회 (테스트 전 시료 조제)
목표 spec.
  - UV RSD(%)≤2%
  - IR RSD(%)≤1%
동일 농도 반복 측정 시 재현성 확인
재현성 : UV – 0.09%, IR – 0.06% (target : 각 ±2% ±1% 이내)
- 시료 변경구간
Quality & Reliability

## Slide 10
2. long-term test [SC-1]
2.1 Test information
Target chemical: H2O2 (5.5 wt%), NH4OH (1.0 wt%)
측정 시간 – 30분 (Long-term)
목표 spec.
  - UV RSD(%)≤2%, IR RSD(%)≤1%
  - H2O2 (wt%) ±0.3%, NH4OH (wt%) ±0.2%
  - Flow rate (ml/min) 15 ml/min (±5 ml)
  - Chemical temperature (℃) ±5 ℃
H2O2 : conc. 5.51 (wt%), RSD(%) 0.31
NH4OH : conc. 1.02 (wt%), RSD(%) 0.51
Chemical temperature: 25.05 ℃
Flow rate: 14.94 (ml/min)
 | NH4OH conc. (wt%) | H2O2 conc. (wt%) | Chem. Temp. (℃) | Flow rate (ml/min)
Target | 1.00 | 5.50 | 25.00 | 15.00
Average | 1.02 | 5.51 | 25.05 | 14.94
STDEV | 0.01 | 0.02 | 0.08 | 0.17
RSD(%) | 0.51 | 0.31 | 0.31 | 1.14
D.I water
SC-1
Quality & Reliability

## Slide 11
New Function (UI)
Extended Touch LCD screen [7.0’’  10.1’’]
UI design version update [UI 1.0  UI 2.0]

## Slide 12
New Function (Auto Calibration & LED life-time)
Auto-calibration for light sources
 | Standard | Advanced (Auto calibration function)
Life-time | 40,000H | 50,000H
Recommended replacement LED | 27,000H | 47,000H
Range of high accuracy | 18,000H | 45,000H
※ The lifetime can vary depending on the user's equipment usage methods.
Auto-calibration applied
Auto-calibration not applied
LED life-time status

## Slide 13
Creativity & Confidence
www.atikorea.com
Thank you