## 98)V0.1보드시험현황
항목 |  | 시료 #1 | 시료 #2 | 시료 #3 | 시료 #4 | 시료 #5 | 시료 #6 | Remark
1. Power | 12V(11.99V) output | 12.11
| 12V(11.99V) ripple (mVpp) | 21.6
| 5V(5.289V) output | 4.995
| 5V(5.289V) ripple (mVpp) | 34
| 5V(A5V) output | 4.994
| 5V(A5V) ripple (mVpp) | 30
| 3.3V output | 3.305
| 3.3V ripple (mVpp) | 26.8
2. F/W upload | IAP image upload | V1.3 | V1.3 | V1.3 | V1.3 | V1.3 | V1.3
| Application image upload | V0.1 | V0.1 | V0.1 | V0.1 | V0.1 | V0.1
3. Board Rework | V/I Convert 저항값 수정 | O | O | O | O | O | O
| BLDC RPM, EN Port Pin swap | O | O | O | O | O | O
| BLDC Buffer OE Pin pull up 추가 | O | O | O | O | O | O
4. UART | UART #1 (DEBUG) | O | O | O | O | O | O
| UART #2 (RS232) | O | O | O | O | O | O
| UART #3 (Panel PC) | O | O | O | O | O | O
| UART #5 (RS485) | O | O | O | O | O | O
| Ethernet | O | O | O | O | O | O
5. Calibration | RTD calibration | O | O | O | O | O | O
| TC calibration | O | O | O | O | O | O
| 4~20mA OUT 0 calibration | O | O | O | O | O | O
| 4~20mA OUT 1 calibration | O | X → O | O | O | O | O
| 4~20mA OUT 2 calibration | O | O | O | O | O | O
| 4~20mA OUT 3 calibration | O | O | O | O | O | O
| 4~20mA IN 0 calibration | X → O | X → O | O | O | O | O
| 4~20mA IN 1 calibration | O | O | O | O | O | O
| 4~20mA IN 2 calibration | O | O | O | O | O | O
| 4~20mA IN 3 calibration | X → O | O | O | O | O | O
6. I/O read/write | Parallel input read (8p) | O | O | O | O | O | O
| Parallel output write (4p) | O | O | O | O | O | O
| Relay contact input read (4p) | O | O | O | O | O | O
| Relay contact output write (8p) | O | O | O | O | O | O
| Sensor input#1 (8p) | O | O | O | O | O | O
| Sensor input#2 (8p) | O | O | O | O | O | O
| 24V output#1 write (1~7ch) | O | O | O | O | O | O
| 24V output#2 write (8~14ch) | O | O | O | O | O | O
| Micro SD card | O | O | O | O | O | O
| BLDC Motor output | O | O | O | O | O | O
| BLDC Motor Input | O | O | O | O | O | O
항목 |  | 시료 #1 | 시료 #2 | 시료 #3 | 시료 #4 | 시료 #5 | 시료 #6 | Remark
MCU Module |  | 조락현사용 | 이동철사용 | PH3 port X | 박순호 사용 | 알파시스템 | PH3 port X
Main Board |  | 조락현사용 | 이동철사용 | 알파시스템 | 박순호 사용 | Reserved | 최기용 사용 | 22.05.17

## 전원체크
<Main Board>
항목 |  | 시료 #1 | 시료 #2 | 시료 #3 | 시료 #4 | 시료 #5 | 시료 #6 | Remark
Power | 12V output, RMS | 12.12 | 12.05 | 12.12 | 12.05 | 12.07 | 12.15
| 12V ripple (mVpp) , max value | 아래참조 | 140 | 142 | 132 | 152 | 138
| 7V output | 6.94 | 6.98 | 6.97 | 6.97 | 7 | 7
| 7V ripple (mVpp) | 아래참조 | 138 | 144 | 136 | 142 | 146
| 5V output(7V→5V) | 4.996 | 4.996 | 4.993 | 5.001 | 4.983 | 5.023
| 5V ripple (mVpp) | 아래참조 | 19.2 | 15.6 | 26.8 | 19.2 | 58
| A5V output(7V→A5V) | 4.993 | 4.972 | 4.973 | 5.017 | 5.001 | 4.977
| A5V ripple (mVpp) | 아래참조 | 18 | 15.2 | 18.8 | 22 | 19.6
| 3.3V output(7V→3.3V) | 3.309 | 3.294 | 3.311 | 3.311 | 3.321 | 3.297
| 3.3V ripple (mVpp) | 아래참조 | 17.6 | 20.8 | 22.8 | 18.4 | 16
시료 #1 각 전원별 ripple 측정 파형
| 1. 12V |  |  | 2. 7V
| 3. 5V |  |  | 4. A5V
| 5. 3.3V |  |  | 6. MCU 25MHz

## Calibration
1. cal 관련 commands
1) cal read/write
| - cal xx | read cal parameter xx (xx = 0 - 255)
| - cal xx yy | write yy to cal parameter xx (xx = 0 - 255, yy = 32-bit unsigned integer)
| - cal save | save current working cal parameters to Flash memory
| - cal show | shows current working cal parameters
| - cal compare | compare cal parameters between Flash and working
2) Flash env table related
| - env cal init | initialize env cal table to default
| - env cal show | shows Flash cal parameters, it can be different with working cal parameters
2. Flash memory 에 저장할 Cal parameter 들
- 이 cal parameter 들은 Flash memory 의 environment 영역에 저장 됨
- board booting 시 여기에서 저장된 cal parameter 들을 읽어서 사용해야 함
Parm No. | Group | Name | Function | Remark | Sample
[5 empty] | #1 | #2 | #3 | #4 | #5,3,3 | #6 | #MCU,Main,Relay
0 | RTD(PT100) sensor calibration data | CAL_RTD0_000 | RTD #1(12-bit) calibration value on 0 °C | default : 269 (range : 242..295) | 253 | 259 | 268 | 268 | 280 | 267
1 |  | CAL_RTD0_100 | RTD #1(12-bit) calibration value on 100 °C | default : 1404 (range : 1263..1544) | 1374 | 1379 | 1407 | 1394 | 1420 | 1399
2 |  | CAL_RTD0_200 | RTD #1(12-bit) calibration value on 200 °C | default : 2481 (range : 2232..2729) | 2436 | 2439 | 2487 | 2457 | 2496 | 2477
3 |  | CAL_RTD1_000 | RTD #2(12-bit) calibration value on 0 °C | default : 269 (range : 242..295)
4 |  | CAL_RTD1_100 | RTD #2(12-bit) calibration value on 100 °C | default : 1404 (range : 1263..1544)
5 |  | CAL_RTD1_200 | RTD #2(12-bit) calibration value on 200 °C | default : 2481 (range : 2232..2729)
6 |  | CAL_RTD_RSV1 | Reserved | default : 0
7 |  | CAL_RTD_RSV2 | Reserved | default : 0
8 | TC sensor calibration data | CAL_TC0_000 | TC #0 (12-bit) calibration value on 0 °C | default : 621 (range : 600..650) | 620 | 622 | 645 | 637 | 604 | 642
9 |  | CAL_TC0_246 | TC #0 (12-bit) calibration value on 246 °C | default : 3227 (range : 3150..3300) | 3237 | 3245 | 3256 | 3247 | 3220 | 3250
10 |  | CAL_TC1_000 | TC #1 (12-bit) calibration value on 0 °C | default : 621 (range : 600..650)
11 |  | CAL_TC1_246 | TC #1 (12-bit) calibration value on 246 °C | default : 3227 (range : 3150..3300)
12 | TC1047 offset calibration data | TC_CJ_OFFSET_SIGN | TC1047 temperature offset sign - TC1047 device is used to detect thermocouple's cold junction temperature. But if there is a offset (temperature difference) between TC1047 and TC's end terminal temperature, use this parameter. - OFFSET = TC1047 temp. - cold junction temp. - OFFSET > 0 : TC_CJ_OFFSET_SIGN = 1 - OFFSET < 0 : TC_CJ_OFFSET_SIGN = 0 | default : 1,
13 |  | TC_CJ_OFFSET_VAL | TC1047 temperature offset value - OFFSET_VAL = ABS(TC1047 temp. - cold junction temp.) - value has x2 format, 1000 = 10.00 | default : 200, ( 2°C, x2 value)
14 | TC1047 | TC_CJ_REPLACE | Fixed constant enviroment temperature this is used when TC1047 cold junction temperature sensor is in abnormal state | default : 3300, (33°C, x2 value)
15 |  | TC_CJ_VALID_MIN | TC1047 temperature valid range min value | default : 500, ( 5°C, x2 value)
16 |  | TC_CJ_VALID_MAX | TC1047 temperature valid range max value | default : 6500, (65°C, x2 value)
17
18 | TC sensor calibration data | CAL_TC_RSV1 | Reserved | default : 0
19 |  | CAL_TC_RSV2 | Reserved | default : 0
20 | ADC_IN0 calibration data | CAL_ADC_IN0_4m | ADC input ch 0 calibrated value for current @4mA | default : 11,000
21 |  | CAL_ADC_IN0_20m | ADC input ch 0 calibrated value for current @20mA | default : 50,000
22 |  | CAL_ADC_IN0_0V | ADC input ch 0 calibrated value for voltage @0V | default : 0 | 37776 | 37912 | 37704 | 38407 | 37898 | 37933
23 |  | CAL_ADC_IN0_480mV | ADC input ch 0 calibrated value for voltage @480mV | default : 50,000 | 60590 | 60287 | 59706 | 60423 | 59691 | 60040
24 | ADC_IN1 calibration data | CAL_ADC_IN1_4m | ADC input ch 1 calibrated value for current @4mA | default : 11,000
25 |  | CAL_ADC_IN1_20m | ADC input ch 1 calibrated value for current @20mA | default : 50,000
26 |  | CAL_ADC_IN1_0V | ADC input ch 1 calibrated value for voltage @0V | default : 0 | 38004 | 37854 | 37691 | 38346 | 38157 | 38182
27 |  | CAL_ADC_IN1_480mV | ADC input ch 1 calibrated value for voltage @480mV | default : 50,000 | 60995 | 59845 | 59692 | 60472 | 60240 | 60526
28 | ADC_IN2 calibration data | CAL_ADC_IN2_4m | ADC input ch 2 calibrated value for current @4mA | default : 11,000
29 |  | CAL_ADC_IN2_20m | ADC input ch 2 calibrated value for current @20mA | default : 50,000
30 |  | CAL_ADC_IN2_0V | ADC input ch 2 calibrated value for voltage @0V | default : 0 | 38633 | 37904 | 37472 | 38152 | 37938 | 37806
31 |  | CAL_ADC_IN2_480mV | ADC input ch 2 calibrated value for voltage @480mV | default : 50,000 | 63918 | 59888 | 59334 | 60111 | 59840 | 59791
32 | ADC_IN3 calibration data | CAL_ADC_IN3_4m | ADC input ch 3 calibrated value for current @4mA | default : 11,000
33 |  | CAL_ADC_IN3_20m | ADC input ch 3 calibrated value for current @20mA | default : 50,000
34 |  | CAL_ADC_IN3_0V | ADC input ch 3 calibrated value for voltage @0V | default : 0 | 37760 | 37635 | 37431 | 37965 | 38169 | 37859
35 |  | CAL_ADC_IN3_480mV | ADC input ch 3 calibrated value for voltage @480mV | default : 50,000 | 60623 | 59496 | 59282 | 59864 | 60149 | 59884
36 | ADC_IN4 calibration data | CAL_ADC_IN4_4m | ADC input ch 4 calibrated value for current @4mA | default : 11,000
37 |  | CAL_ADC_IN4_20m | ADC input ch 4 calibrated value for current @20mA | default : 50,000
38 |  | CAL_ADC_IN4_0V | ADC input ch 4 calibrated value for voltage @0V | default : 0
39 |  | CAL_ADC_IN4_480mV | ADC input ch 4 calibrated value for voltage @480mV | default : 50,000
40 | ADC_IN5 calibration data | CAL_ADC_IN5_4m | ADC input ch 5 calibrated value for current @4mA | default : 11,000
41 |  | CAL_ADC_IN5_20m | ADC input ch 5 calibrated value for current @20mA | default : 50,000
42 |  | CAL_ADC_IN5_0V | ADC input ch 5 calibrated value for voltage @0V | default : 0
43 |  | CAL_ADC_IN5_480mV | ADC input ch 5 calibrated value for voltage @480mV | default : 50,000
44 | ADC_IN6 calibration data | CAL_ADC_IN6_4m | ADC input ch 6 calibrated value for current @4mA | default : 11,000
45 |  | CAL_ADC_IN6_20m | ADC input ch 6 calibrated value for current @20mA | default : 50,000
46 |  | CAL_ADC_IN6_0V | ADC input ch 6 calibrated value for voltage @0V | default : 0
47 |  | CAL_ADC_IN6_480mV | ADC input ch 6 calibrated value for voltage @480mV | default : 50,000
48 | ADC_IN7 calibration data | CAL_ADC_IN7_4m | ADC input ch 7 calibrated value for current @4mA | default : 11,000
49 |  | CAL_ADC_IN7_20m | ADC input ch 7 calibrated value for current @20mA | default : 50,000
50 |  | CAL_ADC_IN7_0V | ADC input ch 7 calibrated value for voltage @0V | default : 0
51 |  | CAL_ADC_IN7_480mV | ADC input ch 7 calibrated value for voltage @480mV | default : 50,000
52 | ADC IN calibration data | CAL_ADC_RSV1 | Reserved | default : 0
53 |  | CAL_ADC_RSV2 | Reserved | default : 0
54
55
56
57
58
59
60 | DAC_OUT0 (external) calibration data | CAL_DAC_OUT0_4m | DAC output ch 0 calibrated value for current @4mA | default 11,000 | 12077 | 12160 | 12038 | 12020 | 12010 | 12134
61 |  | CAL_DAC_OUT0_20m | DAC output ch 0 calibrated value for current @20mA | default 59,000 | 58045 | 60320 | 59330 | 59070 | 59860 | 60180
62 |  | CAL_DAC_OUT0_0V | DAC output ch 0 calibrated value for voltage @0V | default : 9800
63 |  | CAL_DAC_OUT0_10V | DAC output ch 0 calibrated value for voltage @10V | default : 59000
64 | DAC_OUT1 (external) calibration data | CAL_DAC_OUT1_4m | DAC output ch 1 calibrated value for current @4mA | default 11,000 | 11931 | 12060 | 12044 | 11960 | 13490 | 12134
65 |  | CAL_DAC_OUT1_20m | DAC output ch 1 calibrated value for current @20mA | default 59,000 | 56308 | 60300 | 59820 | 59170 | 61930 | 59900
66 |  | CAL_DAC_OUT1_0V | DAC output ch 1 calibrated value for voltage @0V | default : 9800
67 |  | CAL_DAC_OUT1_10V | DAC output ch 1 calibrated value for voltage @10V | default : 59000
68 | DAC_OUT2 (external) calibration data | CAL_DAC_OUT2_4m | DAC output ch 2 calibrated value for current @4mA | default 11,000 | 11966 | 12100 | 11910 | 11900 | 11890 | 12100
69 |  | CAL_DAC_OUT2_20m | DAC output ch 2 calibrated value for current @20mA | default 59,000 | 56781 | 59820 | 58820 | 59060 | 59250 | 60180
70 |  | CAL_DAC_OUT2_0V | DAC output ch 2 calibrated value for voltage @0V | default : 9800
71 |  | CAL_DAC_OUT2_10V | DAC output ch 2 calibrated value for voltage @10V | default : 59000
72 | DAC_OUT3 (external) calibration data | CAL_DAC_OUT3_4m | DAC output ch 3 calibrated value for current @4mA | default 11,000 | 11970 | 12140 | 12016 | 12000 | 11960 | 12080
73 |  | CAL_DAC_OUT3_20m | DAC output ch 3 calibrated value for current @20mA | default 59,000 | 57215 | 59730 | 59110 | 59290 | 59520 | 60390
74 |  | CAL_DAC_OUT3_0V | DAC output ch 3 calibrated value for voltage @0V | default : 9800
75 |  | CAL_DAC_OUT3_10V | DAC output ch 3 calibrated value for voltage @10V | default : 59000
76 | DAC_OUT4 (MCU internal) calibration data | CAL_DAC_OUT4_4m | DAC output ch 4 calibrated value for current @4mA | default : 700
77 |  | CAL_DAC_OUT4_20m | DAC output ch 4 calibrated value for current @20mA | default : 3,800
78 |  | CAL_DAC_OUT4_0V | DAC output ch 4 calibrated value for voltage @0V | default : 600
79 |  | CAL_DAC_OUT4_10V | DAC output ch 4 calibrated value for voltage @10V | default : 3700
80 | DAC_OUT5 (MCU internal) calibration data | CAL_DAC_OUT5_4m | DAC output ch 5 calibrated value for current @4mA | default 700
81 |  | CAL_DAC_OUT5_20m | DAC output ch 5 calibrated value for current @20mA | default 3,800
82 |  | CAL_DAC_OUT5_0V | DAC output ch 5 calibrated value for voltage @0V | default : 600
83 |  | CAL_DAC_OUT5_10V | DAC output ch 5 calibrated value for voltage @10V | default : 3700
84 | DAC_OUT6 (MCU internal) calibration data | CAL_DAC_OUT6_4m | DAC output ch 6 calibrated value for current @4mA | default : 700
85 |  | CAL_DAC_OUT6_20m | DAC output ch 6 calibrated value for current @20mA | default : 3,800
86 |  | CAL_DAC_OUT6_0V | DAC output ch 6 calibrated value for voltage @0V | default : 600
87 |  | CAL_DAC_OUT6_10V | DAC output ch 6 calibrated value for voltage @10V | default : 3700
88 | DAC_OUT7 (MCU internal) calibration data | CAL_DAC_OUT7_4m | DAC output ch 7 calibrated value for current @4mA | default 700
89 |  | CAL_DAC_OUT7_20m | DAC output ch 7 calibrated value for current @20mA | default 3,800
90 |  | CAL_DAC_OUT7_0V | DAC output ch 7 calibrated value for voltage @0V | default : 600
91 |  | CAL_DAC_OUT7_10V | DAC output ch 7 calibrated value for voltage @10V | default : 3700
92 | DAC_OUT calibration data | CAL_DAC_RSV1 | Reserved | default : 0
93 |  | CAL_DAC_RSV2 | Reserved | default : 0
94
95
96
97
98
99
100
101
102
103
104
105
106
107
108
109
110
111
112
113
114
115
116
117
118
119
120
121
122
123
124
125
126
127
128
129
130
131
132
133
134
135
136
137
138
139
140
141
142
143
144
145
146
147
148
149
150
151
152
153
154
155
156
157
158
159
160
161
162
163
164
165
166
167
168
169
170
171
172
173
174
175
176
177
178
179
180
181
182
183
184
185
186
187
188
189
190
191
192
193
194
195
196
197
198
199
200
201
202
203
204
205
206
207
208
209
210
211
212
213
214
215
216
217
218
219
220
221
222
223
224
225
226
227
228
229
230
231
232
233
234
235
236
237
238
239
240
241
242
243
244
245
246
247
248
249
250
251
252
253
254
Note on G43: [스레드 댓글] 사용 중인 버전의 Excel에서 이 스레드 댓글을 읽을 수 있지만 파일을 이후 버전의 Excel에서 열면 편집 내용이 모두 제거됩니다. 자세한 정보: https://go.microsoft.com/fwlink/?linkid=870924. 댓글: LM7322 교체
Note on H43: [스레드 댓글] 사용 중인 버전의 Excel에서 이 스레드 댓글을 읽을 수 있지만 파일을 이후 버전의 Excel에서 열면 편집 내용이 모두 제거됩니다. 자세한 정보: https://go.microsoft.com/fwlink/?linkid=870924. 댓글: LM7322교체
Note on G55: [스레드 댓글] 사용 중인 버전의 Excel에서 이 스레드 댓글을 읽을 수 있지만 파일을 이후 버전의 Excel에서 열면 편집 내용이 모두 제거됩니다. 자세한 정보: https://go.microsoft.com/fwlink/?linkid=870924. 댓글: LM7322 교체
Note on H85: [스레드 댓글] 사용 중인 버전의 Excel에서 이 스레드 댓글을 읽을 수 있지만 파일을 이후 버전의 Excel에서 열면 편집 내용이 모두 제거됩니다. 자세한 정보: https://go.microsoft.com/fwlink/?linkid=870924. 댓글: FET(SI2319)교체