## 98)V0.1보드시험현황
|  | 박순호 책임 | 조락현
|  | SW 테스트 | TEST 중
항목 |  | 시료 #1 | 시료 #2 | 시료 #3 | 시료 #4 | 시료 #5 | 시료 #6 | Remark
1. Power | 12V output | 12.11 | 12.04 | 12.09 | 12.01 | 12.04 | 11.99
| 12V ripple (mVpp) | 38 | 32 | 34 | 36 | 36 | 40
| 5V output | 5.077 | 5.092 | 5.096 | 5.075 | 5.09 | 5.107
| 5V ripple (mVpp) | 24 | 24 | 26 | 22 | 24 | 28
| 3.3V output | 3.293 | 3.31 | 3.286 | 3.297 | 3.305 | 3.291
| 3.3V ripple (mVpp) | 20 | 26 | 24 | 22 | 24 | 26
2. F/W upload | IAP image upload | V0.1 | V0.1 | V0.1 | V0.1 | V0.1 | V0.1
| Application image upload | V0.1 | V0.1 | V0.1 | V0.1 | V0.1 | V0.1
3. Board Rework | GND Jumper pin 추가 | O | O
| LTC2601 SDO Pin 단락 | O | O | O
| Digital output 회로 수정 (Jump) | O | O | O
| SPI Damping 저항 교체 | O | O
| UART4 series 저항 교체 | O | O
4. 통신 | UART #1 (Debug/PC IO Frame) | O | O
| UART #2 (PC Data Frame) | O | O
| UART #3 (APM200i) | O | O
| UART #6 (PLC) | O | O
5. Calibration | 4~20mA Ext DAC OUT 0 calibration | O | O | O
| 4~20mA Ext DAC OUT 1 calibration | O | O | O
| 4~20mA Ext DAC OUT 2 calibration | O | O | O
| 4~20mA Ext DAC OUT 3 calibration | O | O | O
| 4~20mA Ext ADC IN 0 calibration | O | O | O
| 4~20mA Ext ADC IN 1 calibration | O | O | O
| 4~20mA Ext ADC IN 2 calibration | O | O | O
| 4~20mA Ext ADC IN 3 calibration | O | O | O
6. I/O read/write | Parallel input read | O | O | O
| Parallel output write | O | O | O
7. DAC | 4~20mA Ext DAC OUT 0 | O | O | O
| 4~20mA Ext DAC OUT 1 | O | O | O
| 4~20mA Ext DAC OUT 2 | O | O | O
| 4~20mA Ext DAC OUT 3 | O | O | O
8. ADC | 4~20mA Ext ADC IN 0 | O | O | O
| 4~20mA Ext ADC IN 1 | O | O | O
| 4~20mA Ext ADC IN 2 | O | O | O
| 4~20mA Ext ADC IN 3 | O | O | O
9. FAN Alarm | FAN1
FAN2

## 99)HW추가수정사항
HW 추가 수정 필요한 사항
| No. | Issued by | 수정 필요 항목 | Description | 반영 여부 | Remark
Rev0.1 | 1 | 조락현 | LTC2601(DAC) SDO(MISO) Open | Daisy chain용 Data out pin 제거
|  |  | SPI Damping Resistor 값 변경 | 220R → 47R
Digital output 회로 수정
|  |  | UART4 직렬 저항 값 변경 | 10K → 10R (R200, R201)

## 00) Sigma-4000_spec
1. Sigma-4000 Basic Specification

## 01) SW작업방법
Sigma-4000 Main board SW 작업 방법
1. 개발 환경 set-up
1) STM32CubeIDE download 및 설치 한다
- STM32CubeIDE 프로그램은 STM32F407IG MCU 사용시 pin 설정, clock 설정, peripheral 설정 등을 수행한 후 기본 초기화 된 C source 를 generation 해주는 utility 이다.
- 또한 통합 개발 환경을 제공하여, compile, link 등을 수행하여 최종 image 까지 만들어 준다
- www.st.com 에서 STM32CubeIDE S/W download 후 압축 풀고 설치
2) STM32CubeProgrammer download 및 설치 한다
- STM32CubeProgrammer 프로그램은 STM32F407IG MCU 사용시 최초 보드 제작후 System boot 기능을 이용하여 IAP 를 program 하기 위한 tool 이다
2. 개발 source 관리
1) 개발 source 는 svn 에 저장하고 version 관리 한다
2) IAP source 저장 위치
https://yjyoo/svn/iap_fw/trunk
3) Nu-2000 main board source 저장 위치
https://yjyoo/svn/oas_dsp/oas_dsp_fw/trunk
3. Source 개발 순서
1) CubeIDE 실행 --> project 생성
2) MCU resource assign (GPIO, I2C, SPI, timer etc)
3) MCU clock assign
이후는 아래 case 와 동일함
4. 다른 project 를 copy 하여 새로운 project 를 만드는 경우
1) CubeMX 를 먼저 실행하여 pjt 이름을 확인하고 디렉토리 이름을 이것으로 변경한다.
2) CubeIDE 실행한다. 이때 workspace 는 project 가 있는 상위 디렉토리를 선택해야 한다.
3) compile error 발생시 흔히 include path 가 빠진것이므로 이를 추가해 준다
4) 각 pin assign 및 MCU resource 할당 (UART, I2C, SPI 등)
5) clock assign
6) source coding & debuggingh
5. 주의 사항
- 반드시 compiler option 의 optimization 을 off 하고 사용할 것

## 01-1) IAP
1. IAP (In-circuit Application Programming) 관련
1) STM32F407 Boot Configuration 관련
a) STM32F407 은 Boot pin 2ea 의 설정상태에 따라 Booting 동작을 선택할 수 있다.
| Boot 1 | Boot 0 | Boot Operation |  | Remark
| x | 0 | Main Flash memory |  | - IAP 설치되어 살아날때 약 2초간 S/W upgrade 경우인지 판단하여 S/W upgrade 를 하거나 그냥 살아나거나를 판단. - IAP program 이 address 0x800_0000 ~ 0x800_FFFF 를 사용하므로 Application program 은 address 0x801_0000 ~ 0x80F_FFFF 를 사용하여야 한다.
| 0 | 1 | System memory |  | - 최초에 IAP 를 program 하기위해 한 번 사용함. - embedded bootloader 가 내장되어 Flash memory 를 reprogram 하는 용도로 사용됨. - UART1/UART3/USB 를 사용함
| 1 | 1 | Embedded SRAM |  | unused
b) 보드가 최초 조립된 이후 DIP switch 를 System memory boot 로 설정한후 System memory boot 기능을 이용하여 미리 작성해둔 IAP 이미지를 Main Flash memory 에 program 한다.
c) 이후 DIP switch 는 Main Flash memory boot 상태로 설정하며 DIP switch 설정을 다시 변경할 경우는 없다.
d) IAP 가 program 된 후에는 Application S/W upgrade 는 serial port #1 을 이용하여 수행한다.
- 필요한 시리얼 통신 프로그램 : YMODEM 으로 파일 upload 가 가능해야 함. (ex. ExtraPutty 등)
- IAP 를 이용한 upgrade mode 상태로 진입하는 방법 : 보드 reset 후 1.5초 이내에 시리얼 통신 프로그램에서 ESC 버튼을 3회 이상 입력
- 필요한 application S/W format : binary
- serial port 설정 : 115200 baud, 1 stop bit, no parity
2) IAP 제작
a) Github 에 있는 STM32F4 의 IAP source 를 참조하여 제작함
- https://github.com/eziya/STM32F4_HAL_IAP_UART 에서 source downloiad 함
- source 수정 내용은 https://blog.naver.com/eziya76/221552811740 의 설명 참조 함
b) CubeIDE tool 에서 new project 생성함
- new project 생성
- pin 설정 : external clock(bypass mode), USART1 pin 번호 설정,
- clock configuration(HSE, use PLL, 168MHz) 설정
- source generation 함
c) 생성한 new project 에 download 한 파일들 추가 및 변경 내용들 비교하여 반영 한다.
- 제일 중요한 내용 : STM32F407 의 MSP 는 0x20020000 이기 때문에, 원래 코드대로 0x2FFE0000 와 AND 연산하면 0x20020000 이 되기 때문에
정상적으로 Application 으로 Jump 를 하지 못하므로 0x2FFD0000 으로 변경한 부분
2. IAP 를 통한 upgrade 를 위해 Application 에서 반드시 수정해야 하는 부분
1) system_stm32f4xx.c : 아래와 같이 변경할 것
- application 의 시작 번지를 0x0801_0000 으로 변경하는 내용, CubeIDE 설치 버전에 따라 아래와 같이 수정이 다를수 있음
2) STM32F407IGHX.ld : 아래와 같이 수정해야 함
- application 의 시작 번지를 0x0801_0000 으로 변경하는 내용
3. Flash Memory Map
1) IAP : sector 0 에 위치하며 sector 4 ~ 11 이 user application SW 저장 됨
sector 1, 2, 3 은 environment config 저장을 위해 reserve 함, 일단 sector 3 을 config 저장 위치로 사용 함
Flash memory 의 특정 sector 를 환경 변수 저장 공간으로 사용함으로 보드상에 있는 EEPROM 은 사용하지 않음
Address | Flash Sector | Size | Nu-2000 Main controller S/W
80F_FFFF | Sector 11 | 128 K
80E_0000 |  | #7
| Sector 10 | 128 K
80C_0000 |  | #6
| Sector 9 | 128 K
80A_0000 |  | #5
| Sector 8 | 128 K
808_0000 |  | #4
| Sector 7 | 128 K
806_0000 |  | #3
| Sector 6 | 128 K
804_0000 |  | #2
| Sector 5 | 128 K
802_0000 |  | #1
| Sector 4 | 64 K | User App program SW starts from here address : 0x801_0000
801_0000 |  | #1
| Sector 3 | 16 K | Reserved area for environment configure #1
800_C000 |  | #4
| Sector 2 | 16 K | Reserved area for environment configure #2
800_8000 |  | #3
| Sector 1 | 16 K | Reserved area for environment configure #3
800_4000 |  | #2
| Sector 0 | 16 K | IAP program start from 0x800_0000 and jump to user app SW at 0x801_0000
800_0000 |  | #1
Environment data
- Flash sector 3 은 512 byte 씩 32 개의 env_parm group 으로 구성 된다.
- Flash memory 의 특성상 erase 시 모든 bit 가 1 이 되며 1 --> 0 으로 write 는 항상 가능한 특성을 이용함
- 이러한 형태로 구성하는 이유는 Flash erase 횟수를 최소화하기 위함 (Flash erase 가능 횟수 : 10,000 회로 통상적인 Flash memory 의 erase 가능 횟수 100,000 회에 비해 작기 때문)
즉, 1회 erase 후 31 번을 write 할 수 있음으로 마치 310,000 회를 write 하는 효과를 볼 수 있음
- 1 parm = 4 byte 로 구성
- 1 group = 128 개의 parm 으로 구성
- 첫번째 group 은 env table 의 version 및 env_parm group 중 valid 한것이 어느것인지 판단하는 용도로 사용
- 실제 valid 한 group 은 1개만 존재함
- env parm group 30 까지 쓴 다음에 Flash sector erase 수행하고 다시 env group 0 번 부터 write 함
Note : Flash erase/program 관련 변동사항 있음 (IDE 버전에서 변경된것으로 추정됨)
erase/program 시 반드시 HAL_FLASH_Unlock(); 과 HAL_FLASH_Lock(); 을 추가해야 함
program 시에는 매번 write 하기 전에 __HAL_FLASH_CLEAR_FLAG(FLASH_FLAG_EOP ¦ FLASH_FLAG_OPERR ¦ FLASH_FLAG_WRPERR ¦ FLASH_FLAG_PGAERR ¦ FLASH_FLAG_PGSERR ); 가 추가되어야 함
sector #3 |  |  | base_address = 0x0800_C000 | [9 empty] | 저장 횟수 : 126 | 63
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | A | B | C | D | E | F
0x0000 | A | T | I | K | _ | E | N | V | _ | 0 | . | 1 |  |  |  |  | Env version
0x0010 | 0.. 7 | 8.. 15 | 16.. 23 | 24.. 31 | [12 empty] | usage bit flag : 0(used), 1(not used)
0x0020 | [16 empty] | unused
0x0030
0x0040
0x0050
0x0060
0x0070 |  |  |  |  | reserved
0x0080
0x0090
0x00A0
0x00B0
0x00C0
0x00D0
0x00E0
0x00F0
0x0100 | CAL parm 10 |  |  |  | CAL parm 11 |  |  |  | CAL parm 12 |  |  |  | CAL parm 13 |  |  |  | calibration data back-up
0x0110 | CAL parm 14 |  |  |  | CAL parm 15 |  |  |  | CAL parm 16 |  |  |  | CAL parm 17
0x0120 | CAL parm 18 |  |  |  | CAL parm 19 |  |  |  | CAL parm 20 |  |  |  | CAL parm 21
0x0130 | CAL parm 22 |  |  |  | CAL parm 23 |  |  |  | CAL parm 24 |  |  |  | CAL parm 25
0x0140 | CAL parm 26 |  |  |  | CAL parm 27 |  |  |  | CAL parm 28 |  |  |  | CAL parm 29
0x0150 | CAL parm 30 |  |  |  | CAL parm 31 |  |  |  | CAL parm 32 |  |  |  | CAL parm 33
0x0160 | CAL parm 34 |  |  |  | CAL parm 35 |  |  |  | CAL parm 36 |  |  |  | CAL parm 37
0x0170 | CAL parm 38 |  |  |  | CAL parm 39
0x0180 | [16 empty] | unused
0x0190
0x01A0
0x01B0 |  |  |  |  | reserved
0x01C0
0x01D0
0x01E0
0x01F0
0x0200 | parm 0 |  |  |  | parm 1 |  |  |  | parm 2 |  |  |  | parm 3 |  |  |  | parm group 0
0x0210 | parm 4 |  |  |  | parm 5 |  |  |  | parm 6 |  |  |  | parm 7
. . .
0x03F0 | parm 124 |  |  |  | parm 125 |  |  |  | parm 126 |  |  |  | parm 127
0x0400 | parm 0 |  |  |  | parm 1 |  |  |  | parm 2 |  |  |  | parm 3 |  |  |  | parm group 1
. . .
0x05F0 | parm 124 |  |  |  | parm 125 |  |  |  | parm 126 |  |  |  | parm 127
0x0600 | parm 0 |  |  |  | parm 1 |  |  |  | parm 2 |  |  |  | parm 3 |  |  |  | parm group 2
. . .
0x07F0 | parm 124 |  |  |  | parm 125 |  |  |  | parm 126 |  |  |  | parm 127
. . .
0x3E00 | parm 0 |  |  |  | parm 1 |  |  |  | parm 2 |  |  |  | parm 3 |  |  |  | parm group 30
. . .
0x3FF0 | parm 124 |  |  |  | parm 125 |  |  |  | parm 126 |  |  |  | parm 127

## 02) GPIO
Sigma-3000/4000 MCU GPIOs |  |  |  |  | SPI | ADC | DAC
STM32F407IG |  |  |  |  | UART | I2C | GPIO
[5 empty] | SDIO | JTAG | CLK OUT
GPIO_A | Function Block | Signal Name | Direction | Function | ETH_MII
0 | GPIO | --- | IN | unused (internal pull-up)
1 | GPIO | --- | IN | unused (internal pull-up)
2 | UART2 | Tx | OUT | LCD(unused) - 115200, 8, no, 1 stop bit
3 | UART2 | Rx | IN
4 | DAC1 | DAC1 | AOUT | PLC_DAC1 (4~20mA)
5 | DAC2 | DAC2 | AOUT | PLC_DAC2 (4~20mA)
6 | ADC1 | ADC1_IN6 | AIN | ADC_SPARE1
7 | ADC2 | ADC2_IN7 | AIN | ADC_SPARE2
8 | GPIO | --- | IN | unused (internal pull-up)
9 | UART1 | Tx | OUT | DEBUG - 115200, 8, no, 1 stop bit
10 | UART1 | Rx | IN
11 | GPIO | --- | IN | unused (internal pull-up)
12 | GPIO | --- | IN | unused (internal pull-up)
13 | JTAG | SWDIO |  | JTAG
14 | JTAG | SWCLK
15 | JTAG | JTDI
GPIO_B | Function Block | Signal Name | Direction | Function
0 | GPIO | --- | IN | unused (internal pull-up)
1 | GPIO | --- | IN | unused (internal pull-up)
2 | BOOT | BOOT1 | IN | Boot1 : always 0
3 | JTAG | SWO |  | JTAG
4 | JTAG | NJTRST
5 | GPIO | --- | IN | unused (internal pull-up)
6 | GPIO | --- | IN | unused (internal pull-up)
7 | GPIO | --- | IN | unused (internal pull-up)
8 | GPIO | --- | IN | unused (internal pull-up)
9 | GPIO | --- | IN | unused (internal pull-up)
10 | UART3 | Tx | OUT | APM200i - 115200, 8, no, 1 stop bit
11 | UART3 | Rx | IN
12 | GPIO | --- | IN | unused (internal pull-up)
13 | GPIO | --- | IN | unused (internal pull-up)
14 | GPIO | --- | IN | unused (internal pull-up)
15 | GPIO | --- | IN | unused (internal pull-up)
GPIO_C | Function Block | Signal Name | Direction | Function
0 | ADC3 | ADC3_IN10 | AIN | ADC_SPARE3
1 | GPIO | --- | IN | unused (internal pull-up)
2 | GPIO | --- | IN | unused (internal pull-up)
3 | GPIO | --- | IN | unused (internal pull-up)
4 | GPIO | --- | IN | unused (internal pull-up)
5 | GPIO | --- | IN | unused (internal pull-up)
6 | UART6 | RS485_Tx | OUT | PLC controller (RS485) - 9600, 8, no, 1 stop bit
7 | UART6 | RS485_Rx | IN
8 | SDIO | SD_D0 | IN/OUT | micro SD card interface use DMA2
9 | SDIO | SD_D1 | IN/OUT
10 | SDIO | SD_D2 | IN/OUT
11 | SDIO | SD_D3 | IN/OUT
12 | SDIO | SD_CLK | OUT
13 | GPIO | SD_DET | IN | SD card detect - avoid chattering min. 400 msec - 0 : SD card extracted, 1 : SD card inserted - internal pull-up required
14 |  | OSC32 |  | external 32.768KHz clock for RTC
15 |  | OSC32
GPIO_D | Function Block | Signal Name | Direction | Function
0 | GPIO | --- | IN | unused (internal pull-up)
1 | GPIO | --- | IN | unused (internal pull-up)
2 | SDIO | CMD |  | micro SD card interface
3 | GPIO | --- | IN | unused (internal pull-up)
4 | GPIO | --- | IN | unused (internal pull-up)
5 | GPIO | --- | IN | unused (internal pull-up)
6 | GPIO | --- | IN | unused (internal pull-up)
7 | GPIO | DAC_OD_1 | OUT | DAC(XTR111) control (High : On, Low : Off)
8 | GPIO | DAC_OD_2 | OUT
9 | GPIO | DAC_OD_3 | OUT
10 | GPIO | DAC_OD_4 | OUT
11 | GPIO | SPI_Select1 | OUT | SPI Chip slect pin SPI_Select1..4 : Ext DAC(LTC2601)#1..4 SPI_Select5 : Ext ADC(AD7682)
12 | GPIO | SPI_Select2 | OUT
13 | GPIO | SPI_Select3 | OUT
14 | GPIO | SPI_Select4 | OUT
15 | GPIO | SPI_Select5 | OUT
GPIO_E | Function Block | Signal Name | Direction | Function
0 | GPIO | MCU_D_OUT_0 | OUT | Digital output 0
1 | GPIO | MCU_D_OUT_1 | OUT | Digital output 1
2 | GPIO | MCU_D_OUT_2 | OUT | Digital output 2
3 | GPIO | MCU_D_OUT_3 | OUT | Digital output 3
4 | GPIO | MCU_D_OUT_4 | OUT | Digital output 4
5 | GPIO | MCU_D_OUT_5 | OUT | Digital output 5
6 | GPIO | MCU_D_OUT_6 | OUT | Digital output 6
7 | GPIO | MCU_D_OUT_7 | OUT | Digital output 7
8 | GPIO | MCU_D_OUT_8 | OUT | Digital output 8
9 | GPIO | MCU_D_OUT_9 | OUT | Digital output 9
10 | GPIO | MCU_D_OUT_10 | OUT | Digital output 10
11 | GPIO | MCU_D_OUT_11 | OUT | Digital output 11
12 | GPIO | MCU_D_OUT_12 | OUT | Digital output 12
13 | GPIO | MCU_D_OUT_13 | OUT | Digital output 13
14 | GPIO | MCU_D_OUT_14 | OUT | Digital output 14
15 | GPIO | MCU_D_OUT_15 | OUT | Digital output 15
GPIO_F | Function Block | Signal Name | Direction | Function
0 | GPIO | TP6 | IN | unused (internal pull-up)
1 | GPIO | TP7 | IN | unused (internal pull-up)
2 | GPIO | TP9 | IN | unused (internal pull-up)
3 | GPIO | TP10 | IN | unused (internal pull-up)
4 | GPIO | FAN_IN_0 | IN | FAN Alarm
5 | GPIO | FAN_IN_1 | IN
6 | GPIO | --- | IN | unused (internal pull-up)
7 | GPIO | --- | IN | unused (internal pull-up)
8 | GPIO | --- | IN | unused (internal pull-up)
9 | GPIO | --- | IN | unused (internal pull-up)
10 | GPIO | --- | IN | unused (internal pull-up)
11 | GPIO | --- | IN | unused (internal pull-up)
12 | GPIO | --- | IN | unused (internal pull-up)
13 | GPIO | LED_YL | OUT | user LED yellow : 1(on), 0(off)
14 | GPIO | LED_GR | OUT | user LED green : 1(on), 0(off)
15 | GPIO | LED_RED | OUT | user LED red : 1(on), 0(off)
GPIO_G | Function Block | Signal Name | Direction | Function
0 | GPIO | MCU_D_IN_0 | IN | Digital input 0
1 | GPIO | MCU_D_IN_1 | IN | Digital input 1
2 | GPIO | MCU_D_IN_2 | IN | Digital input 2
3 | GPIO | MCU_D_IN_3 | IN | Digital input 3
4 | GPIO | MCU_D_IN_4 | IN | Digital input 4
5 | GPIO | MCU_D_IN_5 | IN | Digital input 5
6 | GPIO | MCU_D_IN_6 | IN | Digital input 6
7 | GPIO | MCU_D_IN_7 | IN | Digital input 7
8 | GPIO | MCU_D_IN_8 | IN | Digital input 8
9 | GPIO | MCU_D_IN_9 | IN | Digital input 9
10 | GPIO | MCU_D_IN_10 | IN | Digital input 10
11 | GPIO | MCU_D_IN_11 | IN | Digital input 11
12 | GPIO | MCU_D_IN_12 | IN | Digital input 12
13 | GPIO | MCU_D_IN_13 | IN | Digital input 13
14 | GPIO | MCU_D_IN_14 | IN | Digital input 14
15 | GPIO | MCU_D_IN_15 | IN | Digital input 15
GPIO_H | Function Block | Signal Name | Direction | Function
0 |  | OSC_IN |  | 25MHz clock input
1 |  | OSC_OUT(TP8)
2 | GPIO | --- | IN | unused (internal pull-up)
3 | GPIO | --- | IN | unused (internal pull-up)
4 | GPIO | --- | IN | unused (internal pull-up)
5 | GPIO | --- | IN | unused (internal pull-up)
6 | GPIO | --- | IN | unused (internal pull-up)
7 | I2C2 | SCL |  | EEPROM I2C I/F (128 Kbyte) --> unused - GPIO input with pull-up
8 | I2C2 | SDA
9 | GPIO | --- | IN | unused (internal pull-up)
10 | GPIO | --- | IN | unused (internal pull-up)
11 | GPIO | --- | IN | unused (internal pull-up)
12 | GPIO | --- | IN | unused (internal pull-up)
13 | GPIO | --- | IN | unused (internal pull-up)
14 | GPIO | --- | IN | unused (internal pull-up)
15 | GPIO | --- | IN | unused (internal pull-up)
GPIO_I | Function Block | Signal Name | Direction | Function
0 | GPIO | --- | IN | unused (internal pull-up)
1 | SPI2 | CLK |  | SPI I/F, used for … ※ Damping resistor : 47R 1) 4ch ADC 1ea (AD7682) 2) 1ch DAC 4ea (LTC2601)
2 | SPI2 | MISO
3 | SPI2 | MOSI
4 | GPIO | TP2 | IN | unused (internal pull-up)
5 | GPIO | TP3 | IN | unused (internal pull-up)
6 | GPIO | TP4 | IN | unused (internal pull-up)
7 | GPIO | TP5 | IN | unused (internal pull-up)
8 | GPIO | PCB Version1 | IN | PCB Issue 00 : Rev0.1, 01 : Rev0.2, 10 : Rev0.3, 11 : Rev0.4
9 | GPIO | PCB Version2 | IN
10 | GPIO | --- | IN | unused (internal pull-up)
11 | GPIO | --- | IN | unused (internal pull-up)
12
13
14
15

## 02-1) Timer
MCU timers
1. Cautions on using timer
- timer start 이전에 반드시 counter/status register 를 clear 하여야 함.
- clear 안하면 timer start 시에 fake overflow interrupt 발생함.
- free running timer 형태로 사용하는 경우, 처음 한번 fake overflow interrupt 발생이 큰 문제 야기하지 않음
- timer 를 one-shot 형태로 start/stop 을 반복하는 형태로 사용하는 경우에는 매번 fake interrupt 에
의해 timer 가 정상 동작하지 않음
2. Timer usage table
| Type | Description | Max clock |  | Prescaler |  |  | Auto Reload |  |  | Psi-3000
|  |  | freq. (MHz) | period (nsec) | range | setting value | freq. | range | setting value | freq.
Independent watchdog | IWDG | prescaler : 8-bit 12-bit down counter | 32 KHz internal RC clk |  | 0 .. 256 | 16 | 2 KHz (0.5 msec) | 0 .. 4096 | 4096 (4096 - 1) | 0.4884 Hz (2.0475 s) | 10 msec main loop에 적용
Window watchdog |  | 7-bit down counter | from main clk | [7 empty] | 미 사용
Timer 1 | Advanced | 16-bit | 168 | 5.95
Timer 2 | General | 32-bit | 84 | 11.9
Timer 3 | General | 16-bit | 84 | 11.9
Timer 4 | General | 16-bit | 84 | 11.9
Timer 5 | General | 32-bit | 84 | 11.9
Timer 6 | Basic | - 16-bit prescaler - 16-bit auto reload upcounter - timer overflow interrupt | 84 | 11.9 | 1 .. 65536 | 84 (84 - 1) | 1 MHz (1 usec) | 1 .. 65536 | 1000 (1000 - 1) | 1 KHz (1 msec) | 1 msec main timer
Timer 7 | Basic | - 16-bit prescaler - 16-bit auto reload upcounter - timer overflow interrupt | 84 | 11.9
Timer 8 | Advanced | 16-bit | 168 | 5.95
Timer 9 | General | 16-bit | 168 | 5.95
Timer 10 | General | 16-bit | 168 | 5.95
Timer 11 | General | 16-bit | 168 | 5.95
Timer 12 | General | 16-bit | 84 | 11.9
Timer 13 | General | 16-bit | 84 | 11.9
Timer 14 | General | 16-bit | 84 | 11.9

## 02-2) SW scheduling
SW scheduling
- minimem loop time : 1msec
- main time base : 10 msec (using timer)
| 0 msec | 10 msec | 20 msec | 30 msec | 40 msec | 50 msec | 60 msec | 70 msec | 80 msec | 90 msec | 100 msec
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 0
@ 1 msec | 0 msec: - UART PC svc
@ 10 msec | - 4~20mA out svc | - 4~20mA out svc | - 4~20mA out svc | - 4~20mA out svc | - 4~20mA out svc | - 4~20mA out svc | - 4~20mA out svc | - 4~20mA out svc | - 4~20mA out svc | - 4~20mA out svc
| - 4~20mA in svc | - 4~20mA in svc | - 4~20mA in svc | - 4~20mA in svc | - 4~20mA in svc | - 4~20mA in svc | - 4~20mA in svc | - 4~20mA in svc | - 4~20mA in svc | - 4~20mA in svc
| - Digital input svc | - Digital input svc | - Digital input svc | - Digital input svc | - Digital input svc | - Digital input svc | - Digital input svc | - Digital input svc | - Digital input svc | - Digital input svc
| - Sys. LED svc | - Sys. LED svc | - Sys. LED svc | - Sys. LED svc | - Sys. LED svc | - Sys. LED svc | - Sys. LED svc | - Sys. LED svc | - Sys. LED svc | - Sys. LED svc
| - Debug svc | - Debug svc | - Debug svc | - Debug svc | - Debug svc | - Debug svc | - Debug svc | - Debug svc | - Debug svc | - Debug svc
| - Monitor svc | - Monitor svc | - Monitor svc | - Monitor svc | - Monitor svc | - Monitor svc | - Monitor svc | - Monitor svc | - Monitor svc | - Monitor svc
@ 100 msec | - monitor svc | - monitor svc | - monitor svc | - monitor svc | - monitor svc | - monitor svc | - monitor svc | - monitor svc | - monitor svc | - monitor svc
@ 200 msec | - sd card ejection svc | - sd card ejection svc | - sd card ejection svc | - sd card ejection svc | - sd card ejection svc | - sd card ejection svc | - sd card ejection svc | - sd card ejection svc | - sd card ejection svc | - sd card ejection svc

## 03) Parm
1. Parameter 관련 commands
1) parameter read/write
| - parm xx | read parameter xx (xx = 0 - 127)
| - parm xx yy | write yy to parameter xx (xx = 0 - 127, yy = 32-bit unsigned integer)
| - parm save | save current working parameters to Flash memory except calibration info.
| - parm show | shows current working parameters
| - parm compare | compare parameters between Flash and working
Note : parameter 10 ~ 39 are reserved for calibration info, thus they will not be saved even "parm save" command entered
2) Flash env table related
| - env init | initialize env table to default except cal_info
| - env save | save current working parameters to Flash memory (same as "parm save")
| - env show | shows Flash parameters, it can be different with working parameters
| - env compare | compare parameters between Flash and working (same as "parm compare")
3) Calibration info related
| - cal load | load calibration info of Flash memory to current working calibration info area
| - cal save | save current working calibration info to Flash memory
| - cal show | shows calibration info
| - cal compare | compare calibration info between Flash and working
2. Flash memory 에 저장할 Parameter 들
- 이 parameter 들은 Flash memory 의 environment 영역에 저장 됨
- board booting 시 여기에서 저장된 parameter 들을 읽어서 사용해야 함
Parm No. | Group | Name | Function | Remark | Rev0.1 Sample
[5 empty] | #1 | #2 | #3 | #4 | #5 | #6
0 |  | mon_onoff | System monitor trace onoff control | default : 1 (monitor on)
1 |  | mon_flag | monitor contents control bit flag - bit 0 : default monitoring (date, temp, adc avg 1sec[0..3]) - bit 1 : add to default monitoring (UV_PD0_adc, UV_PD1_adc, IR_PD0_adc, IR_PD0_adc) - bit 2 : ADC value only (for cal.) - bit 3 : RTD 0..1 raw ADC value - bit 4 : 4-20mA IN(MFC) ADC_avg, current avg (16) - bit 5 : 4-20mA OUT#0-#3 consent_val, ADC_val (32) - bit 6 : MFC in/out monitoring (64) | default : 0x0001
2 |  | mon_interval | monitoring interval based on 1 sec | default : 3 sec (range 1..60)
3 |  | OUT_420_EN | 4-20 mA current output port onoff control --> data : 0(off), 1(on) - bit 0 : 4-20mA port #0 (to PLC #0) - bit 1 : 4-20mA port #1 (to PLC #1) - bit 2 : 4-20mA port #2 (to MFC) - bit 3 : 4-20mA port #3 (reserved) | default : 15 (0x000f)
4 |  | OUT_420_TEST_FLAG | 4-20 mA current output test flag data : 0 - normal operation, (use calculated consentration) 1 - test operation, (use parm 120 .. 123 value) - bit 0 : 4-20mA port #0 (to PLC #0) - bit 1 : 4-20mA port #1 (to PLC #1) - bit 2 : 4-20mA port #2 (to MFC) - bit 3 : 4-20mA port #3 (reserved) | default : 0x0000
5 |  | TRACE_FLAG | other trace onoff control flag - bit 0 : other trace onoff | default : 0x0000
6 |  | NEX_TRACE_FLAG | nextion trace onoff control flag - bit 0 : Nextion trace onoff - bit 1 : Nextion trace time/date onoff | default : 0x0000
7 |  | PID_TC_TRACE_FLAG | temp. PID trace onoff control flag - bit 0 : temp. PID trace onoff - bit 1 : detail temp. PID trace onoff | default : 0x0000
8
9 |  | PASSWORD | password | 0
10 | RTD(PT100) sensor calibration data | CAL_RTD0_000 | RTD #0(16-bit) calibration value on 0 °C | default : 269 (range : 242..295)
11 |  | CAL_RTD0_100 | RTD #0(16-bit) calibration value on 100 °C | default : 1404 (range : 1263..1544)
12 |  | CAL_RTD0_200 | RTD #0(16-bit) calibration value on 200 °C | default : 2481 (range : 2232..2729)
13 |  | CAL_RTD1_000 | RTD #1(16-bit) calibration value on 0 °C | default : 269 (range : 242..295)
14 |  | CAL_RTD1_100 | RTD #1(16-bit) calibration value on 100 °C | default : 1404 (range : 1263..1544)
15 |  | CAL_RTD1_200 | RTD #1(16-bit) calibration value on 200 °C | default : 2481 (range : 2232..2729)
16
17
18
19
20
21
22 | C4_20 IN0 calibration data | CAL_420_IN0_4m | Current 4-20 mA input #0, calibrated 4mA ADC value | default 9000 | 9550 | 9594 | 9579
23 |  | CAL_420_IN0_20m | Current 4-20 mA input #0, calibrated 20mA ADC value | default 50000 | 47682 | 47991 | 47978
24 | C4_20 IN1 calibration data | CAL_420_IN1_0V | Current 4-20 mA input #0, calibrated 4mA ADC value | default 9000 | 9651 | 9526 | 9618
25 |  | CAL_420_IN1_5V | Current 4-20 mA input #0, calibrated 20mA ADC value | default 50000 | 48098 | 47699 | 48097
26 | C4_20 IN2 calibration data | CAL_420_IN2_0V | Current 4-20 mA input #0, calibrated 4mA ADC value | default 9000 | 9575 | 9611 | 9609
27 |  | CAL_420_IN2_5V | Current 4-20 mA input #0, calibrated 20mA ADC value | default 50000 | 47808 | 48073 | 48090
28 | C4_20 IN3 calibration data | CAL_420_IN3_0V | Current 4-20 mA input #0, calibrated 4mA ADC value | default 9000 | 9583 | 9606 | 9656
29 |  | CAL_420_IN3_5V | Current 4-20 mA input #0, calibrated 20mA ADC value | default 50000 | 47869 | 47992 | 48227
30 | C4_20 OUT0 calibration data | CAL_420_OUT0_4m | Current 4-20 mA output #0, calibrated 4mA DAC value | default 12,800 | 11989 | 11957 | 11968
31 |  | CAL_420_OUT0_20m | Current 4-20 mA output #0, calibrated 20mA DAC value | default 64,000 | 59704 | 59605 | 59890
32 | C4_20 OUT1 calibration data | CAL_420_OUT1_4m | Current 4-20 mA output #1, calibrated 4mA DAC value | default 12,800 | 11986 | 11932 | 12048
33 |  | CAL_420_OUT1_20m | Current 4-20 mA output #1, calibrated 20mA DAC value | default 64,000 | 59699 | 59580 | 60170
34 | C4_20 OUT2 calibration data | CAL_420_OUT2_4m | Current 4-20 mA output #2, calibrated 4mA DAC value | default 12,800 | 11963 | 11904 | 12052
35 |  | CAL_420_OUT2_20m | Current 4-20 mA output #2, calibrated 20mA DAC value | default 64,000 | 59680 | 59450 | 60180
36 | C4_20 OUT3 calibration data | CAL_420_OUT3_0V | Current 4-20 mA output #3, calibrated 4mA DAC value | default 12,800 | 12001 | 11897 | 12054
37 |  | CAL_420_OUT3_5V | Current 4-20 mA output #3, calibrated 20mA DAC value | default 64,000 | 59640 | 59520 | 60085
40 | RTD(PT100) sensor temperature limit | RTD0_TEMP_LIMIT_HH | RTD #0 temperature upper error limit | default = 3000, x100 value, 3000 = 30.0
41 |  | RTD0_TEMP_LIMIT_H | RTD #0 temperature upper warning limit | default = 2800, x100 value
42 |  | RTD0_TEMP_LIMIT_L | RTD #0 temperature lower warning limit | default = 2200, x100 value
43 |  | RTD0_TEMP_LIMIT_LL | RTD #0 temperature lower error limit | default = 2000, x100 value
44 |  | RTD1_TEMP_LIMIT_HH | RTD #1 temperature upper error limit | default = 3000, x100 value, 3000 = 30.0
45 |  | RTD1_TEMP_LIMIT_H | RTD #1 temperature upper warning limit | default = 2800, x100 value
46 |  | RTD1_TEMP_LIMIT_L | RTD #1 temperature lower warning limit | default = 2200, x100 value
47 |  | RTD1_TEMP_LIMIT_LL | RTD #1 temperature lower error limit | default = 2000, x100 value
48 |  | RTD2_TEMP_LIMIT_HH | RTD #2 temperature upper error limit | default = 3000, x100 value, 3000 = 30.0
49 |  | RTD2_TEMP_LIMIT_H | RTD #2 temperature upper warning limit | default = 2800, x100 value
50 |  | RTD2_TEMP_LIMIT_L | RTD #2 temperature lower warning limit | default = 2200, x100 value
51 |  | RTD2_TEMP_LIMIT_LL | RTD #2 temperature lower error limit | default = 2000, x100 value
52 |  | RTD3_TEMP_LIMIT_HH | RTD #3 temperature upper error limit | default = 3000, x100 value, 3000 = 30.0
53 |  | RTD3_TEMP_LIMIT_H | RTD #3 temperature upper warning limit | default = 2800, x100 value
54 |  | RTD3_TEMP_LIMIT_L | RTD #3 temperature lower warning limit | default = 2200, x100 value
55 |  | RTD3_TEMP_LIMIT_LL | RTD #3 temperature lower error limit | default = 2000, x100 value
56 | PID temp controller | PID_TC_SET_VAL | pid temp. controller, temp. set value | default = 250, x10 value
57
58
59
60 | UV LED | uv_led_op_mode | UV LED operation mode : 0..3 | default : 1 0: variable current, 1..3 : fixed current mode
61 |  | uv_led_var_cur | UV LED variable current level : 0..100 mA | default : 100 mA only valid when uv_led_op_mode = 0
62 |  | uv_led_on_dur | UV LED on duration (usec) | default 1,000 range (1..65,000)
63
64
65 | IR LED | ir_led_op_mode | IR LED operation mode : 0..3 | default : 1 0: variable current, 1..3 : fixed current mode
66 |  | ir_led_var_cur | IR LED variable current level : 0..200 mA | default : 200 mA only valid when uv_led_op_mode = 0
67 |  | ir_led_on_dur | IR LED on duration (usec) | default 1,000 range (1..65,000)
68
69
70 | UV PD | uv_pd_meas_point | UV PD measure point before LED off (usec) | default 300 range (1..65,000)
71 |  | uv_pd0_limit_HH | UV PD0 valid range low limit | reference PD
72 |  | uv_pd0_limit_LL | UV PD0 valid range high limit
73 |  | uv_pd1_limit_HH | UV PD1 valid range low limit | measure PD
74 |  | uv_pd1_limit_LL | UV PD1 valid range high limit
75 | IR PD | ir_pd_meas_point | IR PD measure point before LED off (usec) | default 300 range (1..65,000)
76 |  | ir_pd0_limit_HH | IR PD0 valid range low limit | reference PD
77 |  | ir_pd0_limit_LL | IR PD0 valid range high limit
78 |  | ir_pd1_limit_HH | IR PD1 valid range low limit | measure PD
79 |  | ir_pd1_limit_LL | IR PD1 valid range high limit
80 | CONCENTRATION #1 | CONCENT1_MAX | concentration #1, measure upper boundary, % | default : 1000, x100 format, 1000 = 10.00
81 |  | CONCENT1_MIN | concentration #1, measure lower boundary, % | default : 0, x100 format, 0 = 0.00
82 |  | CONCENT1_LIMIT_HH | concentration #1, upper error limit, % | default : 800, x100 format, 800 = 8.00
83 |  | CONCENT1_LIMIT_H | concentration #1, upper warning limit, % | default : 700, x100 format, 700 = 7.00
84 |  | CONCENT1_LIMIT_L | concentration #1, lower warning limit, % | default : 200, x100 format, 200 = 2.00
85 |  | CONCENT1_LIMIT_LL | concentration #1, lower error limit, % | default : 100, x100 format, 100 = 1.00
86 |  | CONCENT1_RESERVE
87 | CONCENTRATION #2 | CONCENT2_MAX | concentration #2, measure upper boundary, % | default : 1000, x100 format, 1000 = 10.00
88 |  | CONCENT2_MIN | concentration #2, measure lower boundary, % | default : 0, x100 format, 0 = 0.00
89 |  | CONCENT2_LIMIT_HH | concentration #2, upper error limit, % | default : 800, x100 format, 800 = 8.00
90 |  | CONCENT2_LIMIT_H | concentration #2, upper warning limit, % | default : 700, x100 format, 700 = 7.00
91 |  | CONCENT2_LIMIT_L | concentration #2, lower warning limit, % | default : 200, x100 format, 200 = 2.00
92 |  | CONCENT2_LIMIT_LL | concentration #2, lower error limit, % | default : 100, x100 format, 100 = 1.00
93 |  | CONCENT2_RESERVE
94 | Mass Flow range : 0 .. 50 (mL/min) | MASS_FLOW_SPEED | mass flow speed (to MFC) | default = 2000, x100 value, 2000 = 20.0
95 |  | MASS_FLOW_LIMIT_HH | mass flow upper error limit | default = 4500, x100 value, 5000 = 50.0
96 |  | MASS_FLOW_LIMIT_H | mass flow upper warning limit | default = 4000, x100 value, 5000 = 50.0
97 |  | MASS_FLOW_LIMIT_L | mass flow lower warning limit | default = 1000, x100 value, 1000 = 10.0
98 |  | MASS_FLOW_LIMIT_LL | mass flow lower error limit | default = 500, x100 value, 1000 = 10.0
99 |  | MASS_RESERVE
100 | Reserved 420 port | RESERVED_420_MAX | concentration #2, measure upper boundary, % | default : 10000, x1000 format, 10000 = 10.000
101 |  | RESERVED_420_MIN | concentration #2, measure lower boundary, % | default : 0, x1000 format, 0 = 0.000
102
103
104
105
106
107
108
109
110 |  | PLC_BAUDRATE | PLC : ASCII MODBUS baud rate (4800/9600/19200/38400/57600/115200) | default : 9600
111 |  | RS485_ADDRESS | RS-485 address | range : 1..247, default = 1
112 |  | modbus_parity | ASCII MODBUS parity (even/odd/no parity)
113 |  | modbus_timeout | ASCII MODBUS message timeout (msec) | 통상 1000
114 |  | PARA0_A1_VAL | concentration 0 calculation parameter : A1 | x100 format
115 |  | PARA0_Y0_VAL | concentration 0 calculation parameter : Y0 | x100 format
116 |  | PARA0_T1_VAL | concentration 0 calculation parameter : T1 | x100 format
117 |  | PARA1_A1_VAL | concentration 1 calculation parameter : A1 | x100 format
118 |  | PARA1_Y0_VAL | concentration 1 calculation parameter : Y0 | x100 format
119 |  | PARA1_T1_VAL | concentration 1 calculation parameter : T1 | x100 format
120 |  | TEST_OUT0_420_VAL | 4-20 mA out #0 current test value | default : 0, range : 0 .. 100
121 |  | TEST_OUT1_420_VAL | 4-20 mA out #1 current test value | default : 0, range : 0 .. 100
122 |  | TEST_OUT2_420_VAL | 4-20 mA out #2 current test value | default : 0, range : 0 .. 100
123 |  | TEST_OUT3_420_VAL | 4-20 mA out #3 current test value | default : 0, range : 0 .. 100
124
125
126
127

## 03-1) Calibration
Calibration 필요한 항목
[5 empty] | CH1 |  | CH2 |  | CH3 |  | CH4
No. | Calibration Items | Description | Remark | No | 0mA | 20mA | 0mA | 20mA | 0mA | 20mA | 0mA | 20mA
1 | 4 - 20 mA input | 4 - 20 mA 입력 1 port 에 대해 internal ADC 및 OP-AMP 주변 부품 오차를 |  | 1 | 9582 | 47986 | 9620 | 48100 | 9610 | 48091 | 9657 | 48226
|  | 측정/저장하여 실제 4 - 20 mA 입력이 이루어질 수 있게 한다. |  | 2 | 9580 | 47983 | 9616 | 48103 | 9611 | 48094 | 9655 | 48227
|  | - 정밀한 전류계를 사용하여 측정한 값과 ADC 측정값 사이의 차이를 비교하여 |  | 3 | 9580 | 47980 | 9615 | 48105 | 9609 | 48090 | 9658 | 48226
|  | 정확한 4mA 와 20mA 가 측정되는 ADC 값을 찾는다. |  | 4 | 9580 | 47980 | 9616 | 48096 | 9609 | 48095 | 9655 | 48227
|  |  |  | 5 | 9579 | 47982 | 9618 | 48094 | 9608 | 48095 | 9652 | 48227
|  | - 4 - 20 mA In #1 : external ADC ch1 calibration |  | 6 | 9576 | 47984 | 9619 | 48097 | 9610 | 48087 | 9654 | 48234
|  | - 4 - 20 mA In #1 : external ADC ch2 calibration |  | 7 | 9576 | 47983 | 9617 | 48095 | 9607 | 48090 | 9660 | 48225
|  | - 4 - 20 mA In #1 : external ADC ch3 calibration |  | 8 | 9577 | 47977 | 9617 | 48096 | 9608 | 48091 | 9659 | 48228
|  | - 4 - 20 mA In #1 : external ADC ch4 calibration |  | 9 | 9581 | 47986 | 9616 | 48103 | 9609 | 48094 | 9660 | 48230
2 | 4 - 20 mA output | 4 - 20 mA 출력 4 port 에 대해 DAC 및 XTR111 주변 부품 오차를 |  | 10 | 9580 | 47982 | 9616 | 48094 | 9612 | 48090 | 9657 | 48230
|  | 측정/저장하여 실제 4 - 20 mA 출력이 이루어질 수 있게 한다. |  | 11 | 9581 | 47983 | 9620 | 48094 | 9606 | 48093 | 9659 | 48232
|  | - DAC 제어값을 변경하면서 정밀한 전류계를 사용하여 측정하여 |  | 12 | 9580 | 47980 | 9618 | 48093 | 9607 | 48096 | 9655 | 48230
|  | 정확한 4mA 와 20mA 가 이루어지는 DAC 제어값을 찾는다. |  | 13 | 9580 | 47979 | 9617 | 48097 | 9610 | 48095 | 9653 | 48228
|  |  |  | 14 | 9579 | 47984 | 9618 | 48093 | 9608 | 48096 | 9658 | 48228
|  | - 4 - 20 mA Out #1 : external DAC#1 (LTC2601) calibration |  | 15 | 9576 | 47981 | 9615 | 48092 | 9615 | 48090 | 9655 | 48231
|  | - 4 - 20 mA Out #2 : external DAC#2 (LTC2601) calibration |  | 16 | 9577 | 47982 | 9616 | 48098 | 9609 | 48086 | 9658 | 48227
|  | - 4 - 20 mA Out #3 : external DAC#3 (LTC2601) calibration |  | 17 | 9582 | 47980 | 9617 | 48098 | 9610 | 48094 | 9651 | 48230
|  | - 4 - 20 mA Out #4 : external DAC#4 (LTC2601) calibration |  | 18 | 9584 | 47983 | 9616 | 48100 | 9607 | 48094 | 9659 | 48225
|  |  |  | 19 | 9575 | 47983 | 9615 | 48098 | 9609 | 48097 | 9660 | 48224
|  |  |  | 20 | 9577 | 47980 | 9617 | 48097 | 9608 | 48095 | 9653 | 48231
|  |  |  | 21 | 9579 | 47981 | 9617 | 48098 | 9609 | 48093 | 9658 | 48227
|  |  |  | 22 | 9578 | 47980 | 9614 | 48099 | 9608 | 48092 | 9658 | 48226
|  |  |  | 23 | 9580 | 47979 | 9618 | 48091 | 9608 | 48097 | 9659 | 48229
|  |  |  | 24 | 9582 | 47979 | 9615 | 48098 | 9606 | 48095 | 9659 | 48231
|  |  |  | 25 | 9583 | 47980 | 9613 | 48099 | 9609 | 48091 | 9655 | 48227
|  |  |  | 26 | 9578 | 47979 | 9616 | 48099 | 9610 | 48091 | 9663 | 48227
|  |  |  | 27 | 9579 | 47981 | 9619 | 48093 | 9610 | 48092 | 9656 | 48227
|  |  |  | 28 | 9579 | 47978 | 9618 | 48098 | 9609 | 48092 | 9658 | 48224
|  |  |  | 29 | 9581 | 47976 | 9620 | 48098 | 9609 | 48092 | 9657 | 48225
|  |  |  | 30 | 9577 | 47978 | 9619 | 48101 | 9611 | 48089 | 9653 | 48224
|  |  |  | 31 | 9579 | 47979 | 9617 | 48099 | 9610 | 48090 | 9658 | 48226
|  |  |  | 32 | 9581 | 47977 | 9620 | 48098 | 9610 | 48087 | 9655 | 48222
|  |  |  | 33 | 9582 | 47975 | 9616 | 48098 | 9611 | 48093 | 9655 | 48223
|  |  |  | 34 | 9580 | 47979 | 9620 | 48095 | 9610 | 48097 | 9657 | 48231
|  |  |  | 35 | 9583 | 47980 | 9615 | 48100 | 9612 | 48089 | 9660 | 48230
|  |  |  | 36 | 9587 | 47975 | 9619 | 48100 | 9611 | 48091 | 9659 | 48225
|  |  |  | 37 | 9586 | 47976 | 9620 | 48094 | 9610 | 48089 | 9655 | 48233
|  |  |  | 38 | 9579 | 47979 | 9614 | 48098 | 9612 | 48088 | 9652 | 48225
|  |  |  | 39 | 9577 | 47977 | 9623 | 48092 | 9611 | 48088 | 9656 | 48223
|  |  |  | 40 | 9578 | 47978 | 9619 | 48099 | 9607 | 48088 | 9655 | 48230
|  |  |  | 41 | 9580 | 47980 | 9621 | 48097 | 9610 | 48091 | 9658 | 48229
|  |  |  | 42 | 9576 | 47977 | 9619 | 48096 | 9611 | 48091 | 9656 | 48232
|  |  |  | 43 | 9574 | 47974 | 9616 | 48100 | 9612 | 48093 | 9654 | 48222
|  |  |  | 44 | 9578 | 47982 | 9617 | 48093 | 9611 | 48092 | 9656 | 48227
|  |  |  | 45 | 9580 | 47973 | 9617 | 48101 | 9607 | 48089 | 9657 | 48219
|  |  |  | 46 | 9579 | 47981 | 9612 | 48096 | 9609 | 48092 | 9658 | 48228
|  |  |  | 47 | 9575 | 47980 | 9624 | 48098 | 9609 | 48088 | 9657 | 48228
|  |  |  | 48 | 9575 | 47980 | 9618 | 48099 | 9612 | 48088 | 9653 | 48228
|  |  |  | 49 | 9581 | 47975 | 9614 | 48095 | 9608 | 48090 | 9656 | 48222
|  |  |  | 50 | 9581 | 47974 | 9618 | 48092 | 9607 | 48091 | 9654 | 48222
|  |  |  | 51 | 9582 | 47980 | 9621 | 48093 | 9609 | 48085 | 9658 | 48229
|  |  |  | 52 | 9575 | 47982 | 9619 | 48096 | 9608 | 48090 | 9652 | 48221
|  |  |  | 53 | 9578 | 47969 | 9620 | 48098 | 9606 | 48086 | 9659 | 48231
|  |  |  | 54 | 9579 | 47984 | 9619 | 48096 | 9608 | 48088 | 9656 | 48230
|  |  |  | 55 | 9578 | 47978 | 9618 | 48100 | 9609 | 48082 | 9658 | 48226
|  |  |  | 56 | 9573 | 47977 | 9617 | 48094 | 9607 | 48089 | 9660 | 48228
|  |  |  | 57 | 9579 | 47981 | 9624 | 48094 | 9611 | 48086 | 9654 | 48223
|  |  |  | 58 | 9578 | 47975 | 9621 | 48094 | 9609 | 48086 | 9656 | 48227
|  |  |  | 59 | 9579 | 47978 | 9617 | 48097 | 9610 | 48083 | 9657 | 48228
|  |  |  | 60 | 9578 | 47977 | 9619 | 48096 | 9610 | 48091 | 9655 | 48227
|  |  |  | 61 | 9581 | 47979 | 9619 | 48090 | 9609 | 48088 | 9653 | 48229
|  |  |  | 62 | 9580 | 47982 | 9619 | 48100 | 9611 | 48087 | 9655 | 48229
|  |  |  | 63 | 9581 | 47975 | 9616 | 48097 | 9606 | 48091 | 9655 | 48227
|  |  |  | 64 | 9582 | 47973 | 9618 | 48093 | 9610 | 48093 | 9654 | 48222
|  |  |  | 65 | 9581 | 47976 | 9619 | 48094 | 9608 | 48086 | 9654 | 48230
|  |  |  | 66 | 9578 | 47976 | 9619 | 48096 | 9609 | 48091 | 9658 | 48223
|  |  |  | 67 | 9578 | 47972 | 9615 | 48096 | 9607 | 48091 | 9654 | 48227
|  |  |  | 68 | 9579 | 47976 | 9617 | 48093 | 9607 | 48086 | 9657 | 48225
|  |  |  | 69 | 9582 | 47976 | 9618 | 48092 | 9605 | 48087 | 9655 | 48226
|  |  |  | 70 | 9578 | 47973 | 9622 | 48095 | 9607 | 48089 | 9654 | 48227
|  |  |  | 71 | 9583 | 47974 | 9619 | 48099 | 9606 | 48086 | 9654 | 48225
|  |  |  | 72 | 9580 | 47970 | 9618 | 48093 | 9610 | 48087 | 9654 | 48227
|  |  |  | 73 | 9575 | 47980 | 9615 | 48094 | 9605 | 48088 | 9654 | 48227
|  |  |  | 74 | 9582 | 47974 | 9618 | 48097 | 9607 | 48093 | 9655 | 48227
|  |  |  | 75 | 9583 | 47976 | 9621 | 48096 | 9607 | 48095 | 9655 | 48224
|  |  |  | 76 | 9576 | 47973 | 9618 | 48100 | 9608 | 48090 | 9654 | 48220
|  |  |  | 77 | 9576 | 47977 | 9617 | 48096 | 9606 | 48090 | 9655 | 48229
|  |  |  | 78 | 9579 | 47979 | 9620 | 48102 | 9604 | 48093 | 9652 | 48228
|  |  |  | 79 | 9584 | 47980 | 9618 | 48099 | 9609 | 48087 | 9656 | 48230
|  |  |  | 80 | 9583 | 47969 | 9618 | 48098 | 9607 | 48094 | 9656 | 48228
|  |  |  | 81 | 9582 | 47981 | 9617 | 48095 | 9612 | 48093 | 9654 | 48233
|  |  |  | 82 | 9582 | 47978 | 9618 | 48100 | 9603 | 48086 | 9654 | 48232
|  |  |  | 83 | 9581 | 47973 | 9616 | 48099 | 9612 | 48086 | 9654 | 48228
|  |  |  | 84 | 9578 | 47974 | 9619 | 48105 | 9608 | 48088 | 9656 | 48226
|  |  |  | 85 | 9576 | 47974 | 9619 | 48101 | 9612 | 48087 | 9654 | 48230
|  |  |  | 86 | 9579 | 47976 | 9618 | 48097 | 9608 | 48090 | 9654 | 48228
|  |  |  | 87 | 9581 | 47976 | 9619 | 48096 | 9609 | 48083 | 9654 | 48226
|  |  |  | 88 | 9581 | 47976 | 9616 | 48096 | 9611 | 48093 | 9654 | 48227
|  |  |  | 89 | 9576 | 47973 | 9617 | 48098 | 9609 | 48088 | 9654 | 48228
|  |  |  | 90 | 9581 | 47979 | 9621 | 48099 | 9608 | 48088 | 9656 | 48230
|  |  |  | 91 | 9576 | 47982 | 9616 | 48101 | 9608 | 48091 | 9652 | 48231
|  |  |  | 92 | 9577 | 47979 | 9615 | 48096 | 9606 | 48088 | 9655 | 48227
|  |  |  | 93 | 9579 | 47979 | 9618 | 48096 | 9611 | 48085 | 9652 | 48227
|  |  |  | 94 | 9581 | 47973 | 9618 | 48099 | 9609 | 48089 | 9658 | 48228
|  |  |  | 95 | 9577 | 47978 | 9619 | 48099 | 9609 | 48093 | 9654 | 48229
|  |  |  | 96 | 9579 | 47980 | 9621 | 48099 | 9610 | 48086 | 9658 | 48225
|  |  |  | 97 | 9579 | 47986 | 9617 | 48098 | 9608 | 48089 | 9653 | 48225
|  |  |  | 98 | 9579 | 47980 | 9619 | 48096 | 9609 | 48093 | 9660 | 48222
|  |  |  | 99 | 9577 | 47978 | 9619 | 48097 | 9612 | 48090 | 9652 | 48233
|  |  |  | 100 | 9581 | 47977 | 9613 | 48097 | 9610 | 48088 | 9657 | 48231
|  |  |  | AVG | 9579 | 47978 | 9618 | 48097 | 9609 | 48090 | 9656 | 48227

## 04) SPI
MCU SPI : SPI #2 shares 5 board resources
1. SPI #2 shared resources
- ADC (AD7682, 4 ch, 16-bit) : 4~20mA input
- DAC (LTC2601, 1 ch, 16-bit) : 4~20mA output
- DAC (LTC2601, 1 ch, 16-bit) : 4~20mA output
- DAC (LTC2601, 1 ch, 16-bit) : 4~20mA output
- DAC (LTC2601, 1 ch, 16-bit) : 4~20mA output
2. AD7682 SPI timing format
1) AD7682 setting & timing
- 4 ch single ended configuration : 현재 회로 그대로 사용
- 4 ch unipolar differencial configuration : 여러 개의 external PD 가 있으며 각각의 출력이 differencial 이므로 사용 불가능 함
| - voltage reference setting : internal reference 4.096V | [11 empty] | Consider internal 4.096V voltage reference by checking ADC variation
if possible, we can eliminate AD8605 for ADC power regulation
- low pass filter : full bandwidth or 1/4 bandwidth
2) AD7682 CFG register
| bit | 13 |  | 12 |  | 11 |  | 10 |  | 9 |  | 8 |  | 7 |  | 6 |  | 5 |  | 4 |  | 3 |  | 2 |  | 1 |  | 0
|  | CFG |  | INCC |  | INCC |  | INCC |  | Inx |  | Inx |  | Inx |  | BW |  | REF |  | REF |  | REF |  | SEQ |  | SEQ |  | RB
|  | 1 |  | 1 |  | 1 |  | 1 |  | 0 | [5 empty] | 1 |  | 0 |  | 0 |  | 1 |  | 0 |  | 0 |  | 1
|  | CFG upd |  | ch configuration (unipolar, GND ref) | [5 empty] | ch selection 0 0 0 : ch 0 0 0 1 : ch 1 0 1 0 : ch 2 0 1 1 : ch 3 | [5 empty] | full BW |  | ref. vtg selection 0 0 1 : int. ref. vtg | [5 empty] | sequencer 0 0 : disable |  |  |  | CFG do not read
- Note 1 : CFG word must be shifted left 2 bits, that means SPI_data[15:2] = CFG[13:0], SPI_data[1:0] = 0
- Note 2 : At power-up, the CFG register is undefined and two dummy conversions are required to update the register
3) MCU SPI settings
- CPOL = CPHA = 0, --> SCK idle state is 0
- MCU SPI block clock freq. = 42 MHz, prescaler = 2 --> SPI_CLK = 21 MHz
4) Timing
- to get 10 valid ADC sampled data, MCU must access ADC 12 times, because it takes time for proper CFG takes place.
first and second 2 ADC values must be ignored because they are garbage data caused by channel changing
- minimum 4 usec time is required between each AD conversion.
300 usec
CNV
| ADC |  | ADC ch0 read 12 times | [8 empty] | ADC ch1 read 12 times | [7 empty] | parm | ADC ch2 read 12 times | [8 empty] | ADC ch3 read 12 times | [8 empty] | ADC ch0 read 12 times
greater than 4usec
[14 empty] | channel #0 ADC read | [30 empty] | channel #1 ADC read
CNV
| ADC |  | ADC ch0 read #0 | [5 empty] | ADC ch0 read #1 | [5 empty] | ADC ch0 read #2 | [5 empty] | ADC ch0 read #3 | [5 empty] | ADC ch0 read #4 | [8 empty] | ADC ch0 read #11 | [6 empty] | ADC ch1 read #0 | [5 empty] | ADC ch1 read #1
|  |  | - ch = 0 | [5 empty] | - ch = 0 | [5 empty] | - ch = 0 | [5 empty] | - ch = 0 | [5 empty] | - ch = 0 | [8 empty] | - ch = 0 | [6 empty] | - ch = 1 | [5 empty] | - ch = 1
|  |  | - data = ch3 | [5 empty] | - data = ch0(-1) | [5 empty] | - data = ch0(0) | [5 empty] | - data = ch0(1) | [5 empty] | - data = ch0(2) | [8 empty] | - data = ch0(9) | [6 empty] | - data = ch0 | [5 empty] | - data = ch1(-1)
MCU | AD7682
GPIO | CNV
SPI_CLK | CLK
MOSI | DIN |  | CFG13 |  | CFG12 |  | CFG11 |  | CFG10 |  | CFG9 |  | CFG8 |  | CFG7 |  | CFG6 |  | CFG5 |  | CFG4 |  | CFG3 |  | CFG2 |  | CFG1 |  | CFG0 |  | x |  | x
MISO | SDO |  | AD15 |  | AD14 |  | AD13 |  | AD12 |  | AD11 |  | AD10 |  | AD9 |  | AD8 |  | AD7 |  | AD6 |  | AD5 |  | AD4 |  | AD3 |  | AD2 |  | AD1 |  | AD0
3. LTC2601 SPI related
- LTC2601 : 1 channel 16-bit DAC
| - output volage range : |  |  | 1) 0 ~ 4.096 V (chip internal reference 4.096V) | [13 empty] | resolution = 4.096 V / 65536 = 0.0625 mV
|  |  |  | 2) 0 ~ 2.500 V (chip internal reference 2.500V) | [13 empty] | resolution = 2.500 V / 65536 = 0.0381 mV
|  |  |  | 3) 0 ~ 5.000 V (external Vcc reference) | [13 empty] | resolution = 5.000 V / 65536 = 0.0763 mV
1) LTC2601 setting
| - initial stage : |  | 1) write CMD = 0 0 1 1 , data = 0x000 (DAC output 0 V)
| - normal operation : |  | 1) write CMD = 0 0 1 1 , data = 0xnnnn
2) LTC2630 timing
100 msec
CSB/LD
| DAC |  | ext. DAC write | [23 empty] | ext. DAC write
CSB/LD
| DAC |  | ext. DAC ch0 write | [5 empty] | ext. DAC ch1 write | [5 empty] | ext. DAC ch2 write | [5 empty] | ext. DAC ch3 write
MCU | LTC2601
GPIO | CSB/LD
SPI_CLK | SCK |  |  | 1 |  | 2 |  | 3 |  | 4 |  | 5 |  | 6 |  | 7 |  | 8 |  | 9 |  | 10 |  | 11 |  | 12 |  | 13 |  | 14 |  | 15 |  | 16 |  | 17 |  | 18 |  | 19 |  | 20 |  | 21 |  | 22 |  | 23 |  | 24
MOSI | SDI |  | C3 |  | C2 |  | C1 |  | C0 |  | 0 |  | 0 |  | 0 |  | 0 |  | D15 |  | D14 |  | D13 |  | D12 |  | D11 |  | D10 |  | D9 |  | D8 |  | D7 |  | D6 |  | D5 |  | D4 |  | D3 |  | D2 |  | D1 |  | D0
MISO | - |  |  |  | CMD | [7 empty] | unused | [20 empty] | DAC data
CMD
|  |  | 0 | 0 | 0 | 0 |  | Write to Input Register
|  |  | 0 | 0 | 0 | 1 |  | Update (Power up) DAC Register
|  |  | 0 | 0 | 1 | 1 |  | Write to and Update (Power up) DAC Register | [14 empty] | Use this command to change DAC output voltage
|  |  | 0 | 1 | 0 | 0 |  | Power down
|  |  | 1 | 1 | 1 | 1 |  | No operation
※ SDO의 경우 Daisy chain용으로 일반 SPI통신용으로 사용 불가

## 05) UART
Sigma-3000/4000 MCU UART Assignment
1. UART 종류 및 구성
UART No. | FW name | Function | setting 관련 | Remark
UART 1 | TPID_DEBUG | DEBUG 및 PC IO Frame 공용 포트 - system monitoring, FW upgrade - 115200 baud, 8 data bit, no parity, 1 stop bit | - interrupt driven | isolated RS-232
UART 2 | TPID_PC | PC Data Frame - 115200 baud, 8 data bit, no parity, 1 stop bit | - interrupt driven | non-isolated RS-232
UART 3 | TPID_APM200i | APM200i 와 연결 - - 115200 baud, 8 data bit, no parity, 1 stop bit | - interrupt driven | isolated RS-232
UART 6 | TPID_PLC | Reserved - - 115200 baud, 8 data bit, no parity, 1 stop bit | - interrupt driven | isolated RS-232

## 05-1) SW IO Protocol
Sigma-3000/4000 SW IO Protocol
1. Protocol 정의
2. Protocol 설계
1) Main Loop (sw_io_comm_service)
→ 1ms 마다 sw_io_comm_service() 함수를 호출
→ UART2(PC)의 Queue에 있는 Char 1Byte를 읽어 (char)ch에 저장한다.
→ ch가 0이 아닐 경우(입력이 있을 경우) sw_io_data_service(ch)를 호출

## 05-2) APM200i Protocol
Sigma-3000/4000 MCU UART Assignment
1. Protocol 정의
2. Protocol 설계
1) Main Loop (ap_comm_service)
→ 1ms 마다 ap_comm_service() 함수를 호출
→ UART3(APM200i)의 Queue에 있는 Char 1Byte를 읽어 (char)ch에 저장한다. | Byte | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50
→ ch가 0이 아닐 경우(입력이 있을 경우) ap_data_service(ch)를 호출 | 24Byte | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | CR | LF
| 48Byte | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | CR | LF
2) Data 처리 (ap_data_service)
→ UART3(APM200i)에 입력이 있을 경우 ap_data_sevice(ch) 함수 호출
→ Buffer overflow 검사 후 문제 없을 시 APRxDataBuf에 (char)ch 저장 (APDataIndex count 증가)
→ 이후 위 단계를 반복하여 (char)ch를 APRxDataBuf에 저장하고, is_end_flag_ap() 함수를 반복 호출하여 end signal(=[CR][LF])가 입력되었는지 확인
→ is_end_flag_ap() 함수에서 end signal이 감지되면 지금까지 입력된 데이터의 길이를 검사 (24byte or 48byte)
(데이터의 길이가 24bit or 48bit가 아닐 경우 [Wrong IO Frame length] Error massage와 함께 SD Card에 Log 기록
→ 데이터의 유효성이 확인되면 APTxDataBuf로 APRxDataBuf의 데이터를 복사 한 후 PC로 TxBuf를 송출
→ Data 송출 후 Tx Buf, Rx Buf 및 Index 모두 초기화
APM200i
Sigma-4000
Main Board
Panel PC
PLC
공급설비
RS232
IO Control
RS232
RS232

## 06) Digital IO
Digital input/output
- APM200i와 공급설비(PLC) 사이에서 Start signal과 Relay signal을 전달하는 기능
- GPIO를 통한 단순한 Read/Write 기능
- Hirose社의 26pin IDC Connector를 사용하고, 8ea Port별로 Isolation을 구성

## 07)Power
Circuit (1 port) | Current(mA)
| 24V | 12V | 5V | 3.3V
Ref | 1 | 1.84 | 3.8419 | 3.8419
| 0.9995 | 1.8391 | 3.84 | 3.84
Idle | 46.3 | 85.192 | 177.8809 | 68 | Idle
| 0 @24V | 0 @12V | 131.13 @5V | 68 @3.3V | 73.3375
Isolation RS232 | 1 | 1.84 | 3.8419 | -
GPIO output | 1.1 | 2.024 | 4.2261 | 6
System LED | 0.12526 | 0.23048 | 0.48125 | 0.7 | 5.8182
DAC(20mA) | 20 | - | - | -
| 63.3 | 116.472 | 243.1935 | 353.7361
1.5192
0.3